규 격 서

HSK NO.	Item No.	Description	Unit	Quantity
		Urethero-Renoscope Set(요관경)	set	1

A. Features

1. Atraumatic distal end

2. Small diameter

3. Central instrument channel 5 Fr.

4. Easy to operate

5. Enhanced user safety

B. Specifications

1. URS

1) Distal tip: 8 Fr.

2) Instrument Sheath: 9.5 Fr. conical, 1 step, 9.5 - 13.5 Fr.

3) Working channel: 6Fr. for use with instrument up to 5 Fr.

4) Length: 43 cm

5) Telescope: Karl Storz Fiber Optic System, direction of view 6˚

6)  Eyepiece : angled, rigid

7)  Instrument : Length: 60 cm,  OD: 5 Fr.

2. Lithoclast

1) Pneumatic Power Source : 0 - 8 bar continuously adjustable

2) Supply Pressure: 8 bar medical compressed air

3) Pulse Mode: Single or multi pulses digitally selected

4) Pulse Rate Control: 50 - 600 pulse/min.

5) Isolation : Internally transformer isolation

6) Sterilization : Handpiece - Chemical or gas

Probe - Autoclave, chanical or gas

7) Dimension : Weight - 4.7 Kg., 306× 255× 118 mm

Hand Piece : 190× 20 mm

C. Consist of

1. URS, long Type  :  1 ea

2. Grasping Forceps 5Fr : 1 ea

3. Biopsy Forceps 5Fr : 1 ea

4. Lithoclast : 1ea

D. Remark

1.  1 years warranty