제품규격서

품목번호	품 명	규 격	단 위	수 량	제조국
	초음파치료기	Intelect Mobile Ultrasound	SET	2	미국

A. 특징

1. 대형 LCD를 통해 치료상태 및 조건 등을 쉽게 파악 할 수 있다.

2. ㄱ자형태의 초음파 도자(Applicator)는 인체공학적으로 설계되어 장시간 사용시 편리하다.

3. 1개의 Applicator(5㎠)로도 본체에서 간단한 조작으로 치료중에도 1MHz와 3MHz를 손쉽게 전환 사용이 가능하다.

4. 치료부위에 따라 치료 도자의 종류를 다양하게(option: 1㎠, 2㎠, 10㎠)선택 사용 할 수 있다.

5. Head Warm 기능을 내장하여 빠르게 치료도자를 따뜻하게 하여 줌으로 쾌적한 치료를	제공한다.

6. 전도 및 치료상태를 알리는 Contact signal -LED indicator 및 일시 멈춤(Pause)기능이 있다.

7. Duty Cycle조정을 10%, 20%, 50%, Continuous등으로 선택할 수 있다.

8. BNR은 5:1을 기준으로 설정되었다.

9. 질환별 치료 기억모드(Clinical Indication Parameters)와 치료자 선택 기억모드 방법이 가능하다.

10. 유선사용 및 배터리방식의 선택적 사용이 가능하다. (배터리: option)

B. 사양

1. 주파수: 1MHz & 3.3MHz

2. 전원: 100 - 240HVC, 50-60Hz

3. 강도설정표시: W/㎠(Intensity) / Watts (Power)

4. Duty Cycle: 10%, 20%, 50%, Continuous

5. Pulse Frequency: 16Hz, 48Hz, 100Hz

6. Pulse Duration

1msec, +/-20% (10% Duty Cycle)

2msec, +/-20% (20% Duty Cycle)

5msec, +/-20% (50% Duty Cycle)

7. Output Accuracy: +/-20% above 10% of max

8. Amplitude

0 to 2.5W/㎠ in Continuous mode,

0-3W/㎠ in Pulsed modes.

9. Power: 0 watt to 10 watts

10. Effective Radating Area: 4.0㎠+/-1.0㎠

11. Maximum Treatment Time: 30 minutes

12. Standard Weight (with applicator and base): 2.3kg

13. Dimensions: H(16.3cm), W(28.8cm)(23.9cm), D(32.8cm)

14. BNR(초음파 순수도) 5:1

C. 구성

- Main Unit .......................................................... 1EA

- 5㎠ Ultrasound Applicator 1 & 3MHz ................... 1EA

- Ultrasound Transmission Gel 9oz bottle ............... 1EA

- Power Supply cable ............................................ 1EA

- Mobile Cart(Domestic) ......................................... 1EA

- Manual Operation CD ......................................... 1EA

D. 비고

- 품질보증 기간은 납품일로 부터 2년 이내 이다.

- 제품의 설치는 사용장소에서 납품자의 책임하에 완료하며, 별도의 비용이 청구되지 않는다.

- 사용자 교육은 2주 이내에 실시하도록 한다.