제품규격서(B)

H.S.K. NO.	ITEM NO.	DESCRIPTION	Unit	Q'ty
9018.90-9010		Operating Light (LED type)	set	1

A. Features

1. Second generation LED: smaller & more compact, with superior light field yet lowerpower consumption

2. High performance LEDs, arranged into 16 "light engines"

3. Long life time of LEDs

4. Infrared free light with cool light, minimum heat to operator and operation field

5. The new multi-lens concept presents a first-rate shadow control plus integrated deep-cavity illumination

6. Trailblazing design combines ergonomics & technology, bringing an optimal working atmosphere

7. Touch-screen operated panel provides intuitive operation

8. Install light at ceiling plate and tube

9. Optimum color with no color silhouettes because the light-mixing takes place where the light arises

10. Variable color temperature form 3800K to 4800K

11. VariLux offers to choose between a circular or an oval light field

12. Depth of illumination provides excellent support and flexibility in any surgical situation, >129cm

13. Easy to replace the light engine individually

14. Easy cleaning and disinfection – no “edges”

15. No motor mechanism is necessary for focusing and changing the light field diameter - Service is not necessary

B. Specification

1. Light intensity at 1m working distance

- 1st light : Minimum : 160,000lux

2. Lamp type : LED

3. Color Temperature : 3800k-4800k, variable

4. Lamp life : 25,000

5. CRI : 95

6. Focusing light field size at 1m distance : 23-35cm

7. Depth of illumination : (L1+L2) : 129CM

8. Focusing : Two different lens systems and a state-of-the-art software

9. LED Replacement : light engine individually

C. Consist of

1. LED Operating Lights (V16) : 1set

2. Sterilizable handle : 1ea

3. Wall control panel (Included dimmer) : 1ea

4. Suspension tube & Arm : 1ea

5. Ceiling plate : 1ea

6. Ceiling cover : 1ea

7. Ceiling tube wired : 1ea

8. Operating & Service manual : 1bk

D. Remark

1.  Two years Warranty after the performance test.